## NI 9201 Detailed Specifications

Module 9201

Signal Type Voltage

Channels 8

Sample Rate  500 kS/s

Resolution (bits) 12

| The following specifications are typical for the range –40 to 70 °C unless otherwise noted. All voltages are relative to COM unless otherwise noted. The specifications are the same for the NI 9201 and the NI 9221 unless otherwise noted. |
| --- |
| Input Characteristics |
| Number of channels |  | 8 analog input channels |
| ADC resolution |  | 12 bits |
| Type of ADC |  | Successive approximation register (SAR) |

| Sample rate (aggregate) |
| --- |
| Module | Maximum Sample Rate (R Series Expansion Chassis) | Maximum Sample Rate (All Other Chassis) |
| NI 9201, single channel | 475 kS/s | 800 kS/s |
| NI 9201, scanning | 475 kS/s | 500 kS/s |
| NI 9221 | 475 kS/s | 800 kS/s |

| Input range |
| --- |
| NI 9201 |  | ±10 V |
| NI 9221 |  | ± 60 V |

| Operating voltage ranges 1 |
| --- |
| Module | Measurement Voltage, Channel-to-COM | Maximum Voltage, Channel-to-Earth Ground or COM-to-Earth Ground |
|  | Min (V) | Typ (V) | Max (V) | Screw Terminal | DSUB |
| NI 9201 | ±10.3 | ±10.53 | ±10.8 | 250 Vrms | ±60 VDC |
| NI 9221 | ±61.4 | ±62.50 | ±63.8 |  |  |

| Overvoltage protection (channel-to-COM) |  | ±100 V |
| --- | --- | --- |

| NI 9201 accuracy (excludes noise) |
| --- |
| Measurement Conditions | Percent of Reading (Gain Error) | Percent of Range* (Offset Error) |
| Calibrated typ (25 °C, ±5 °C) | ±0.04% | ±0.07% |
| Calibrated max (–40 to 70 °C) | ±0.25% | ±0.25% |
| Uncalibrated typ (25 °C, ±5 °C) | ±0.26% | ±0.46% |
| Uncalibrated max (–40 to 70 °C) | ±0.67% | ±1.25% |
| * Range equals 10.53 V |

| NI 9221 accuracy (excludes noise) |
| --- |
| Measurement Conditions | Percent of Reading (Gain Error) | Percent of Range* (Offset Error) |
| Calibrated typ (25 °C, ±5 °C) | ±0.04% | ±0.07% |
| Calibrated max (–40 to 70 °C) | ±0.25% | ±0.25% |
| Uncalibrated typ (25 °C, ±5 °C) | ±0.26% | ±0.43% |
| Uncalibrated max (–40 to 70 °C) | ±0.67% | ±1.06% |
| * Range equals 62.50 V |

| Stability |
| --- |
| Gain drift |  | ±34 ppm/°C |
| Offset drift |
| NI 9201 |  | ±100 μV/°C |
| NI 9221 |  | ±580 μV/°C |
| Input bandwidth (–3 dB) |
| NI 9201 |  | 690 kHz min |
| NI 9221 |  | 950 kHz min |
| Input impedance |
| Resistance |  | 1 MΩ |
| Capacitance |  | 5 pF |
| Input noise (code-centered) |
| RMS |  | 0.7 LSBrms |
| Peak-to-peak |  | 5 LSB |
| No missing codes |  | 12 bits |
| DNL |  | –0.9 to 1.5 LSB |
| INL |  | ±1.5 LSB |
| Crosstalk (at 10 kHz) |  | –75 dB |
| Settling time (to 1 LSB) |
| NI 9201 |  | 2 μs |
| NI 9221 |  | 1.25 μs |
| MTBF |  | 1,092,512 hours at 25 °C; Bellcore Issue 2, Method 1, Case 3, Limited Part Stress Method |
| Note Contact NI for Bellcore MTBF specifications at other temperatures or for MIL-HDBK-217F specifications. |
| Power Requirements |
| Power consumption from chassis |
| Active mode |  | 1 W max |
| Sleep mode |  | 1 mW max |
| Thermal dissipation (at 70 °C) |
| Active mode |  | 1 W max |
| Sleep mode |  | 32 mW max |
| Physical Characteristics |
| If you need to clean the module, wipe it with a dry towel. Note For two-dimensional drawings and three-dimensional models of the C Series module and connectors, visit ni.com/dimensions and search by module number. |
| Screw-terminal wiring |  | 12 to 24 AWG copper conductor wire with 10 mm (0.39 in.) of insulation stripped from the end |
| Torque for screw terminals |  | 0.5 to 0.6 N · m (4.4 to 5.3 lb · in.) |
| Ferrules |  | 0.25 mm2 to 2.5 mm2 |
| Weight |
| NI 9201/9221 with screw terminal |  | 150 g (5.3 oz) |
| NI 9201/9221 with DSUB |  | 145 g (5.1 oz) |
| Safety |
| Safety Voltages |
| Connect only voltages that are within the following limits. |
| Channel-to-COM |  | ±60 VDC max |
| NI 9201/9221 with Screw Terminal Isolation Voltages |
| Channel-to-channel |  | None |
| Channel-to-earth ground |
| Continuous |  | 250 Vrms, Measurement Category II |
| Withstand |  | 2,300 Vrms, verified by a 5 s dielectric withstand test |
| Measurement Category II is for measurements performed on circuits directly connected to the electrical distribution system. This category refers to local-level electrical distribution, such as that provided by a standard wall outlet, for example, 115 V for U.S. or 230 V for Europe. Caution Do not connect the NI 9201/9221 with screw terminal to signals or use for measurements within Measurement Categories III or IV. |
| NI 9201/9221 with DSUB Isolation Voltages |
| Channel-to-channel |  | None |
| Channel-to-earth ground |
| Continuous |  | 60 VDC, Measurement Category I |
| Withstand |  | 1,000 Vrms, verified by a 5 s dielectric withstand test |
| Measurement Category I is for measurements performed on circuits not directly connected to the electrical distribution system referred to as MAINS voltage. MAINS is a hazardous live electrical supply system that powers equipment. This category is for measurements of voltages from specially protected secondary circuits. Such voltage measurements include signal levels, special equipment, limited-energy parts of equipment, circuits powered by regulated low-voltage sources, and electronics. Caution Do not connect the NI 9201/9221 with DSUB to signals or use for measurements within Measurement Categories II, III, or IV. |
| Hazardous Locations |
| U.S. (UL) |  | Class I, Division 2, Groups A, B, C, D, T4; Class I, Zone 2, AEx nC IIC T4 |
| Canada (C-UL) |  | Class I, Division 2, Groups A, B, C, D, T4; Class I, Zone 2, Ex nC IIC T4 |
| Europe (DEMKO) |  | EEx nC IIC T4 |
| Safety Standards |
| This product is designed to meet the requirements of the following standards of safety for electrical equipment for measurement, control, and laboratory use: IEC 61010-1, EN 61010-1 UL 61010-1, CSA 61010-1 Note For UL and other safety certifications, refer to the product label or the Online Product Certification section. |
| Electromagnetic Compatibility |
| This product meets the requirements of the following EMC standards for electrical equipment for measurement, control, and laboratory use: EN 61326 (IEC 61326): Class A emissions; Industrial immunity EN 55011 (CISPR 11): Group 1, Class A emissions AS/NZS CISPR 11: Group 1, Class A emissions FCC 47 CFR Part 15B: Class A emissions ICES-001: Class A emissions Note For the standards applied to assess the EMC of this product, refer to the Online Product Certification section. Note For EMC compliance, operate this device with double-shielded cables. |
| CE Compliance |
| This product meets the essential requirements of applicable European Directives, as amended for CE marking, as follows: 2006/95/EC; Low-Voltage Directive (safety) 2004/108/EC; Electromagnetic Compatibility Directive (EMC) Note For the standards applied to assess the EMC of this product, refer to the Online Product Certification section. |
| Online Product Certification |
| Refer to the product Declaration of Conformity (DoC) for additional regulatory compliance information. To obtain product certifications and the DoC for this product, visit ni.com/certification, search by module number or product line, and click the appropriate link in the Certification column. |
| Shock and Vibration |
| To meet these specifications, you must panel mount the system. If you are using the NI 9201/9221 with screw terminal, you also must either affix ferrules to the ends of the terminal wires or use the NI 9932 backshell kit to protect the connections. |
| Operating vibration |
| Random (IEC 60068-2-64) |  | 5 grms, 10 to 500 Hz |
| Sinusoidal (IEC 60068-2-6) |  | 5 g, 10 to 500 Hz |
| Operating shock (IEC 60068-2-27) |  | 30 g, 11 ms half sine, 50 g, 3 ms half sine, 18 shocks at 6 orientations |
| Environmental |
| National Instruments C Series modules are intended for indoor use only but may be used outdoors if installed in a suitable enclosure. Refer to the manual for the chassis you are using for more information about meeting these specifications. |
| Operating temperature (IEC 60068-2-1, IEC 60068-2-2) |  | –40 to 70 °C |
| Storage temperature (IEC 60068-2-1, IEC 60068-2-2) |  | –40 to 85 °C |
| Ingress protection |  | IP 40 |
| Operating humidity (IEC 60068-2-56) |  | 10 to 90% RH, noncondensing |
| Storage humidity (IEC 60068-2-56) |  | 5 to 95% RH, noncondensing |
| Maximum altitude |  | 2,000 m |
| Pollution Degree |  | 2 |
| Environmental Management |
| National Instruments is committed to designing and manufacturing products in an environmentally responsible manner. NI recognizes that eliminating certain hazardous substances from our products is beneficial not only to the environment but also to NI customers. For additional environmental information, refer to the NI and the Environment Web page at ni.com/environment. This page contains the environmental regulations and directives with which NI complies, as well as other environmental information not included in this document. |
| Waste Electrical and Electronic Equipment (WEEE) |
| EU Customers At the end of their life cycle, all products must be sent to a WEEE recycling center. For more information about WEEE recycling centers and National Instruments WEEE initiatives, visit ni.com/environment/weee.htm. |
| Calibration |
| You can obtain the calibration certificate for this device at ni.com/calibration. |
| Calibration interval |  | 1 year |

1 Refer to the Safety Guidelines section in the  NI 9201/9221 Operating Instructions and Specifications for more information about safe operating voltages.

## Pinouts/Front Panel Connections

NI 9201/9221 Terminal and Pin Assignments