NI 9472

6 to 30 V

8 channels

750 mA/ch

100 µs

Screw-terminal, D-Sub

Sourcing digital output

## NI 9472 Detailed Specifications

Output CharacteristicsNI 9472 Power RequirementsNI 9474 Power RequirementsPhysical CharacteristicsSafetyEnvironmentalShock and VibrationElectromagnetic CompatibilityCE ComplianceOnline Product CertificationEnvironmental ManagementWaste Electrical and Electronic Equipment (WEEE)

The following specifications are typical for the range –40 to 70 °C unless otherwise noted. All voltages are relative to COM unless otherwise noted. The specifications are the same for the NI 9472 and the NI 9474 unless otherwise noted.

| Output Characteristics |
| --- |
| Number of channels |  | 8 digital output channels |
| Output type |  | Sourcing |
| Power-on output state |  | Channels off |
| External power supply voltage range (Vsup) |
| NI 9472 |  | 6–30 VDC |
| NI 9474 |  | 5–30 VDC |
| Output impedance (R0) |
| Typical |  | 0.07 Ω |
| Maximum |  | 0.13 Ω |
| Continuous output current (I0), per channel |
| NI 9472 |  | 0.75 A max |
| NI 9474 |  | 1 A max |
| Output voltage (V0) |  | Vsup – (I0 · R0 ) |
| I/O protection |
| Voltage |  | 30 VDC max |
| Reversed voltage |  | None |

| Short-circuit behavior |
| --- |
| Current | Channel Behavior | Module Protection |
| 0 to 1 A | Channel does not trip | Module is not damaged |
| 1 to 6 A | Channel does not trip | Module may be damaged |
| 6 to 13 A | Channel may trip | Module may be damaged |
| >13 A | Channel trips | Module is not damaged |

| Short-circuit trip time |  | 10 μs at 13 A |
| --- | --- | --- |
| Output delay time (full load) |
| NI 9472 |  | 100 μs max |
| NI 9474 |  | 1 μs max |
| MTBF |
| NI 9472 |  | 1,113,301 hours at 25 °C; Bellcore Issue 2, Method 1, Case 3, Limited Part Stress Method |
| NI 9474 |  | 479,889 hours at 25 °C; Bellcore Issue 2, Method 1, Case 3, Limited Part Stress Method |
| Note Contact NI for Bellcore MTBF specifications at other temperatures or for MIL-HDBK-217F specifications. |
| NI 9472 Power Requirements |
| Power consumption from chassis |
| Active mode |  | 230 mW max |
| Sleep mode |  | 0.4 mW max |
| Thermal dissipation (at 70 °C) |
| Active mode |  | 1.5 W max |
| Sleep mode |  | 55 mW max |
| NI 9474 Power Requirements |
| Power consumption from chassis |
| Active mode |  | 660 mW max |
| Sleep mode |  | 0.6 mW max |
| Thermal dissipation (at 70 °C) |
| Active mode |  | 1.5 W max |
| Sleep mode |  | 0.6 mW max |
| Physical Characteristics |
| If you need to clean the module, wipe it with a dry towel. |
| Screw-terminal wiring |  | 12 to 24 AWG copper conductor wire with 10 mm (0.39 in.) of insulation stripped from the end |
| Torque for screw terminals |  | 0.5 to 0.6 N · m (4.4 to 5.3 lb · in.) |
| Ferrules |  | 0.25 mm2 to 2.5 mm2 |
| Weight |
| NI 9472/9474 with screw terminal |  | 150 g (5.3 oz) |
| NI 9472 with DSUB |  | 145 g (5.1 oz) |
| Safety |
| NI 9472/9474 with Screw Terminal Safety Voltages |
| Connect only voltages that are within the following limits. |
| Channel-to-COM |  | 30 VDC max |
| Isolation |
| Channel-to-channel |  | None |
| Channel-to-earth ground |
| Continuous |  | 250 Vrms, Measurement Category II |
| Withstand |  | 2,300 Vrms, verified by a 5 s dielectric withstand test |
| Measurement Category II is for measurements performed on circuits directly connected to the electrical distribution system. This category refers to local-level electrical distribution, such as that provided by a standard wall outlet (e.g., 115 V for U.S. or 230 V for Europe). Examples of Measurement Category II are measurements performed on household appliances, portable tools, and similar products. Caution Do not connect the NI 9472/9474 with screw terminal to signals or use for measurements within Measurement Categories III or IV. |
| NI 9472 with DSUB Safety Voltages |
| Connect only voltages that are within the following limits. |
| Channel-to-COM |  | 30 VDC max |
| Isolation |
| Channel-to-channel |  | None |
| Channel-to-earth ground |
| Continuous |  | 60 VDC, Measurement Category I |
| Withstand |  | 1,000 Vrms, verified by a 5 s dielectric withstand test |
| Measurement Category I is for measurements performed on circuits not directly connected to the electrical distribution system referred to as MAINS 1 voltage. This category is for measurements of voltages from specially protected secondary circuits. Such voltage measurements include signal levels, special equipment, limited-energy parts of equipment, circuits powered by regulated low-voltage sources, and electronics. Caution Do not connect the NI 9472 with DSUB to signals or use for measurements within Measurement Categories II, III, or IV. |
| Safety Standards |
| This product is designed to meet the requirements of the following standards of safety for electrical equipment for measurement, control, and laboratory use: IEC 61010-1, EN 61010-1 UL 61010-1, CSA 61010-1 Note For UL and other safety certifications, refer to the product label or the Online Product Certification section. |
| Hazardous Locations |
| U.S. (UL) |  | Class I, Division 2, Groups A, B, C, D, T4; Class I, Zone 2, AEx nC IIC T4 |
| Canada (C-UL) |  | Class I, Division 2, Groups A, B, C, D, T4; Class I, Zone 2, Ex nC IIC T4 |
| Europe (DEMKO) |  | EEx nC IIC T4 |
| Environmental |
| National Instruments C Series modules are intended for indoor use only but may be used outdoors if installed in a suitable enclosure. Refer to the manual for the chassis you are using for more information about meeting these specifications. |
| Operating temperature (IEC 60068-2-1, IEC 60068-2-2) |  | –40 to 70 °C |
| Storage temperature (IEC 60068-2-1, IEC 60068-2-2) |  | –40 to 85 °C |
| Ingress protection |  | IP 40 |
| Operating humidity (IEC 60068-2-56) |  | 10 to 90% RH, noncondensing |
| Storage humidity (IEC 60068-2-56) |  | 5 to 95% RH, noncondensing |
| Maximum altitude |  | 2,000 m |
| Pollution Degree (IEC 60664) |  | 2 |
| Shock and Vibration |
| To meet these specifications, you must panel mount the system. If you are using the NI 9472/9474 with screw terminal, you also must either affix ferrules to the ends of the terminal wires or use the NI 9932 backshell kit to protect the connections. |
| Operating vibration |
| Random (IEC 60068-2-64) |  | 5 grms, 10 to 500 Hz |
| Sinusoidal (IEC 60068-2-6) |  | 5 g, 10 to 500 Hz |
| Operating shock (IEC 60068-2-27) |  | 30 g, 11 ms half sine, 50 g, 3 ms half sine, 18 shocks at 6 orientations |
| Electromagnetic Compatibility |
| This product is designed to meet the requirements of the following standards of EMC for electrical equipment for measurement, control, and laboratory use: EN 61326 EMC requirements; Industrial Immunity EN 55011 Emissions; Group 1, Class A CE, C-Tick, ICES, and FCC Part 15 Emissions; Class A Note For EMC compliance, operate this device with double-shielded cables. |
| CE Compliance |
| This product meets the essential requirements of applicable European Directives, as amended for CE marking, as follows: 2006/95/EC; Low-Voltage Directive (safety) 2004/108/EC; Electromagnetic Compatibility Directive (EMC) Note For the standards applied to assess the EMC of this product, refer to the Online Product Certification section. |
| Online Product Certification |
| Refer to the product Declaration of Conformity (DoC) for additional regulatory compliance information. To obtain product certifications and the DoC for this product, visit ni.com/certification, search by module number or product line, and click the appropriate link in the Certification column. |
| Environmental Management |
| National Instruments is committed to designing and manufacturing products in an environmentally responsible manner. NI recognizes that eliminating certain hazardous substances from our products is beneficial not only to the environment but also to NI customers. For additional environmental information, refer to the NI and the Environment Web page at ni.com/environment. This page contains the environmental regulations and directives with which NI complies, as well as other environmental information not included in this document. |
| Waste Electrical and Electronic Equipment (WEEE) |
| EU Customers At the end of their life cycle, all products must be sent to a WEEE recycling center. For more information about WEEE recycling centers and National Instruments WEEE initiatives, visit ni.com/environment/weee.htm. |

1 MAINS is defined as the (hazardous live) electrical supply system to which equipment is designed to be connected for the purpose of powering the equipment. Suitably rated measuring circuits may be connected to the MAINS for measuring purposes.

## Pinouts/Front Panel Connections

NI 9472/9474 with Screw Terminal

NI 9472 with DSUB